## Test_descrip
EMG Analysis with in-house circuit
Placed circuit & belt-mounted unit in copper box which was grounded to power supply ground
FREQUENCY RESPONSE TESTS
Source Input: | [10 empty] | Input Couple: |  |  | Signals:
| Swept Sine; Log Sweep; Sweep Down (high f to low f); Log Resolution | [10 empty] | DC |  | Grounded Signals (instead of floating)
Frequency: Start @ 1Hz; Span 4 decades; Range: Auto Up & Down
|  | Attenuator Gains: |  |  |  | Rt (kΩ) | R1(kΩ) | Attenuator Gain
[5 empty] | All tests | 0.058 | 100.2 | 0.00057851
[5 empty] | Test #7 | 1.53 | 100.2 | 0.01504
[12 empty] | Low cut-off |  |  | High cut-off
Test # | Saved Name | Plot | Ch | Gain set | Input Ch1 | Source Level (mVpk) | Atten. Gain | Input Level (mVpk) | Output level (Vpp) | Output / Input Level (dB) | Gain Pass (dB) | Fc (Hz) | Gc (dB) | Pc (deg) | Fc (kHz) | Gc (dB) | Pc (deg) | Phase Shift (deg) | Comments
1 | EMGCH1 | Yes | 1 | 10k | After attenuator | 200 | 0.00057851 | 0.1157 | 4.62 @275Hz | 86.0055 | 86.25 | 22.39 | 83.23 | -265 | 2.042 | 83.24 | -463 | 198 | Sinusoids distorted at peaks; appears to be a low freq wave over top of sines
2 | EMGCH2 | Yes | 2 | 10k | After attenuator | 200 | 0.00057851 | 0.1157 | 4.875 @235Hz | 86.4721 | 86.26 | 23.44 | 83.23 | -265 | 2.089 | 83.25 | -463 | 198
3 | EMGCH3 | No | 3 | 10k | After attenuator | 200 | 0.00057851 | 0.1157 | 4.906 @240Hz | 86.5272 | 86.25 | 23.17 | 83.28 | -627 | 2.042 | 83.28 | -822 | 195
4 | EMGCH4 | No | 4 | 10k | After attenuator | 200 | 0.00057851 | 0.1157 | 4.5 @211Hz | 85.7769 | 86.3 | 23.17 | 83.27 | -265 | 2.089 | 83.25 | -464 | 199
5 | EMGCH4A | No | 4 | 10k | After attenuator | 300 | 0.00057851 | 0.17355 | 6.93 @242Hz | 86.0055 | 86.25 | 22.65 | 83.24 | -265 | 2.089 | 83.27 | -464 | 199 | More sinusoidal looking; When @400mVpk input the output was 8Vpp & the 'hi' saturation warning came on
6 | EMGCH4B | Yes | 4 | 1k | After attenuator | 3500 | 0.00057851 | 2.0248 | 7.5 @252Hz | 65.3531 | 66 | 21.88 | 62.9 | -263 | 2.089 | 63.07 | -463 | 200 | More sinusoidal looking
7 | EMGCH4C | Yes | 4 | 100 | After attenuator | 1200 | 0.01504 | 18.0478 | 7.0 @263Hz | 45.7529 | 46 | 22.39 | 43 | -263 | 2.065 | 43.02 | -461 | 198 | Changed Trimpot to Rt = 1.53 kOhm, compared with all other tests at Rt = 58 Ohm
8 | CH4OUTB | No | 4 | 10k | Differential Output B | 300 | 0.00057851 | 0.17355 | 6.75 @252Hz | 85.7769 | 86.3 | 22.65 | 83.27 | -265 | 2.089 | 83.27 | -104 | 161
9 | CH4OUTA | No | 4 | 10k | Differential Output A | 300 | 0.00057851 | 0.17355 | 6.75 @256Hz | 85.7769 | 86.3 | 22.9 | 83.3 | -86 | 2.065 | 83.3 | -283 | 197
COHERENCE TESTS
Input: Random Noise |  |  |  | Averaging: Stable (mean); No = 32
Test # | Saved Name | Plot | Ch | Gain set | Input Ch1 | Source Level (mVpk) | Atten. Gain | Input Level (mVpk) | Comments
10 | COHCH4 | Yes | 4 | 10k | After attenuator | 200 | 0.00057851 | 0.1157 | There was low coherence in the passband; Sudden drops at 100Hz & 1kHz
11 | COHCH4A | No | 4 | 10k | After attenuator | 350 | 0.00057851 | 0.20248 | The coherence was closer to 1.0 in the passband, however still a weird sudden drop @ 10Hz.
Description of Range - Auto Up & Down
For each input channel, the range should be set to ensure the input signal stays in the top half.
Auto Up & Down' allows automatic adjustment to the input range to ensure that the signal level is at least half-scale but not greater than full scale & this is
allowed in either direction.
Description of Averaging - Stable
Often called RMS averaging, it takes the arithmetic mean.
Description of Grounded Signals
Channels are grounded to the Dynamic Signal Analyzer chassis ground.

## Gains
Expected Gain for various gain setting
**Electrodes have gain of 10, but it is included in the overall amplification setting
Overall gain setting | Expected Gain (dB) | Recorded Gain (dB) [taking into account differential factor of 2] | Actual Recorded Gain (dB) | Expected Gain (dB) [taking into account differential factor of 2]
100 | 40 | 39.9794 | 46 | 46.0206
1000 | 60 | 59.9794 | 66 | 66.0206
10000 | 80 | 79.9794 | 86 | 86.0206
| 20 x log(Gain) | Rec. Gain - [20 x log(2)] | From plots | 20 x log(Gain x 2)